Runaway electron probe
Jaroslav Čeřovský

---
Basic description
- Probe mounted on reciprocating mechanism
- Probe consists from: fluorescing crystal covered by a graphite housing
- RE probe was developed and used at TEXTOR
- From measured signals distribution function of RE can be found (under some assumptions)
- In fact similar diagnostic as Cherenkov detector

---
First design of probe
- Probe consist from: scintillating crystals (YSO-Y2SiO5:Ce), optical fibres, photomultipliers,
- 5 mm thick graphite housing, second scintillator behind 8 mm of tungsten
- Probe consist from: scintillating crystals (YSO-Y2SiO5:Ce), optical fibres, photomultipliers,
- 5 mm thick graphite housing, second scintillator behind 8 mm of tungsten
Graphite housing shields from ambient light and low energy particles ( electrons: > 2.5 MeV)

---
Problems of first design
- Piece of tungsten separates runaways from el. direction from particles from ion direction, so second scintillator measures positrons and background
- However, runaways are scattered at the top plate of graphite and second scintillator sees also runaways
- First channel of detector strongly saturated during runaway discharges

---
Second design of probe
- Aim resolve runaways with energies between 5-30 MeV (energy resolution of 3-4 MeV)
- Similar YSO crystals used, scintillators shielded by different thickness of tungsten filters – ranging from 0 mm to 6 mm (top and i. Side – 6-10 mm)

---
Second design of the probe
- Simulations shows (GEANT4 code): sensitivity of probe to electrons is of two orders of magnitude higher than to x-rays
- Probe insensitive to neutrons

---
Third design of probe
Stainless steel filters used instead of tungsten ones → difference in critical energy, less x-ray created in new probe → energy resolution increased

---
Absolute calibration
Performed on electron linac (ELBE)
Experimental setup: Faraday cup, integrated current transformer, Lanex screens

---
Deconvolution of recorded signals
Spectral analysis based on:
- Minimum detectable energy of the crystals
- Same light output of electrons

---
Experiments with probe – setup